# Flexion

Remote Movement Screening for Telehealth

Technical & Clinical Specification — Prototype Stage · Revised July 2026 (literature-review update)

# 1. Overview

Telehealth removed physical observation from the musculoskeletal (MSK) exam. Flexion restores it: during a video visit, or as a between-visit check-in, the patient records a short guided movement routine on their phone or laptop camera. On-device pose estimation converts that video into joint angles, range of motion (ROM), left/right symmetry, and stability scores. Results are attached to the visit as structured data — numbers a provider can trend over time, not just a video to rewatch.

This revision incorporates a structured review of 20 papers on telehealth MSK assessment and markerless/single-camera motion capture (see literature_research_on_the_app.docx and literature_citations_on_app.csv). The review's central finding — that measurement accuracy is highly specific to movement type, joint, and plane of motion, not a single number that applies to "pose estimation" generally — is now reflected directly in the product: every score Flexion produces carries an explicit evidence tier (High / Moderate / Exploratory), and the movement battery itself changed to favor the domains the literature supports best. A working prototype implementing this is referenced throughout (flexion-prototype/).

# 2. System Architecture

Four components, deliberately separated so the clinically-sensitive part (scoring) can be validated and versioned independently of the product surface around it:

| Component | Responsibility | Prototype status |
| --- | --- | --- |
| Capture client | Guides the patient through the movement routine; checks camera framing before scoring starts; runs pose estimation and scoring in the browser so raw video never has to leave the device by default. | Implemented (index.html) |
| Scoring engine | Pure functions mapping pose landmarks → joint angles → ROM / symmetry / stability scores, each tagged with an evidence-graded confidence tier and citation. Framework-agnostic so the exact same math can run client-side or server-side for re-scoring/audits. | Implemented (js/scoring.js), unit tested (20 tests) |
| Session API / store | Persists sessions, links them to patient + provider + visit, exposes them to the dashboard and to EHR export. Not built in the prototype — the prototype substitutes browser localStorage. | Stubbed (localStorage) |
| Provider dashboard | Trends scores over time per patient/movement, surfaces flags and confidence tiers, and lets a provider push a session into the chart as a documented exam finding. | Implemented (dashboard.html) |

Processing pose estimation on-device (rather than uploading raw video to a server) is both a latency and a privacy decision: it avoids transmitting or storing patient video by default, which simplifies the HIPAA data-flow story considerably (see Section 6). The production system should preserve this property — video stays local; only derived numeric scores and, optionally, short annotated clips the provider explicitly requests, leave the device.

# 3. Pose Detection & Scoring Methodology

The prototype uses MediaPipe's Pose Landmarker (BlazePose), a single-camera, markerless model that outputs 33 body landmarks per frame with normalized x/y/z coordinates and a visibility confidence score, running as WebAssembly in the browser. This is consistent with the broader single-camera markerless capture literature reviewed for this product, which the sections below draw on directly.

## 3.1 Joint angle calculation

Every joint angle is computed the same way: given three landmarks A (proximal), B (the joint vertex), and C (distal), the interior angle at B is the angle between vectors BA and BC, via the standard dot-product formula, clamped to [−1, 1] before acos to avoid floating-point domain errors. This is the same geometric definition a goniometer measures physically.

## 3.2 What the literature actually supports, movement by movement

The reviewed literature is specific about where single-camera markerless capture is and isn't accurate yet, and Flexion's evidence tiers are built directly from these findings rather than a general claim of "AI accuracy."

| Domain | Literature finding | Flexion's response |
| --- | --- | --- |
| Sagittal-plane hip/knee kinematics, near-static positions | Mean error <1.1° and limits of agreement <5° vs. lab-based marker systems, holding up even in space-constrained clinical settings (hallways, not just labs). | Squat depth (captured at the near-static bottom of the movement) is scored High confidence. |
| Spatiotemporal gait parameters (step count, cadence, timing) | Good-to-excellent inter-rater reliability and concurrent validity, ICC 0.81–0.98. | Walk-in-place cadence/step count scored High confidence; step-height asymmetry (a kinematic proxy) scored lower. |
| Dynamic transitional movements (e.g., sit-to-stand) | Joint-angle measurements during sit-to-stand show wide limits of agreement (12–20°) even in studies where other markerless measures were reliable. | Sit-to-stand is scored by rep count and total time — the established Five-Times-Sit-to-Stand functional/fall-risk metric — and deliberately does not report a knee-angle depth. |
| Simple (single-plane) shoulder movement: flexion, abduction | Good-to-excellent validity for simple planar movements; complex/rotational shoulder movement is inconsistent across studies. | Arm-raise module measures flexion/abduction only, scored Moderate confidence; not claimed for rotational mobility. |
| Balance / functional movement scoring | The least evidence-covered domain identified in the review across clinical need, validation data, and standardization. | Balance stability score scored Exploratory — presented as a within-patient trend, not an absolute measurement, pending pilot validation. |
| Smartphone/consumer-camera conditions generally | Overall gait RMSE of 5.8°, peak error 11.3°, in a smartphone-based single-camera validation — above the 2–5° threshold generally considered clinically desirable, though comparable to some IMU/commercial markerless systems. | Framing/tracking-quality check added before scoring starts; every report surfaces its evidence tier so a provider isn't given false precision. |

## 3.3 Evidence tiers in the product

Every RoutineResult (Section 4) now carries an evidence object — tier (high / moderate / exploratory), a plain-language note, and a citation — generated from a single EVIDENCE table in js/scoring.js so the confidence framing shown to a provider in the dashboard, in the patient-facing report, and in the exported FHIR Observation's note field are all guaranteed to match. This directly addresses what the literature review flagged as the biggest structural gap: not whether remote movement capture works at all, but whether it produces standardized, trustworthy, decision-grade data rather than an undifferentiated number.

## 3.4 Camera setup and framing

Multiple reviewed papers connect measurement accuracy to camera setup and full-body visibility, and separately call for standardized virtual-exam protocols (patient positioning, camera framing, environment) rather than ad hoc video calls. The capture client now enforces a brief framing check — full-body landmark visibility must hold for roughly 1.2 seconds — before the timed routine begins, and displays a live tracking-quality indicator throughout the session.

# 4. Data Model

Core entities for the session API (not yet built — the prototype substitutes localStorage on the capture/dashboard clients):

| Entity | Key fields |
| --- | --- |
| Patient | id, external EHR patient ID (for matching), name, DOB, assigned provider(s) |
| Provider | id, external EHR provider ID, name, practice |
| Session | id, patientId, providerId, visitId (optional link to a telehealth encounter), capturedAt, deviceInfo, routineVersion |
| RoutineResult | sessionId, movement, metrics (JSON: angles/ROM/asymmetry/stability per movement type), evidence (tier, note, citation), flags[], rawLandmarkRef (optional, for audit/re-scoring) |
| ScoringModelVersion | id, releasedAt, changelog — every RoutineResult records which scoring model version produced it, so historical trends stay comparable after model updates |

Recording rawLandmarkRef (not raw video) alongside each result is what makes retrospective re-scoring, audits, and A/B testing of new scoring models possible without ever needing to re-transmit or re-store patient video.

# 5. EHR Integration

Each RoutineResult is designed to map 1:1 onto a FHIR Observation resource — the prototype already emits this shape, including the evidence tier as an Observation.note (see buildReport() in js/capture.js). Production integration is standard SMART on FHIR: OAuth2 app launch from the EHR, Observation.write scope, category="exam", with the movement name and structured metrics in the Observation payload and component fields for individual angle/ROM values so they're queryable, not just a JSON blob.

## 5.1 Billing context (RTM)

The 2022 CPT code set for Remote Therapeutic Monitoring (RTM) — 98975, 98976, 98977, 98980, 98981 — explicitly covers musculoskeletal system status monitoring, and is the natural billing pathway for a between-visit functional monitoring product like this one. 98977 covers device supply/data-transmission for MSK monitoring (16–30 days in a 30-day period; 2026 national average payment ~$40), and 98980/98981 cover the clinician's monthly treatment-management time reviewing that data. CMS's CY2026 Physician Fee Schedule further expanded RTM, including a new lower-threshold device code (2–15 days of data vs. the prior 16-day minimum) — which directly lowers the bar for how often a patient needs to complete a Flexion routine to support billing.

# 6. Security & Compliance

- PHI minimization by design: video is processed on-device and not persisted or transmitted by default; only derived scores (numbers, not imagery) leave the device unless a provider explicitly requests a clip.

- HIPAA: as a vendor handling PHI (scores tied to patient identity) on behalf of covered-entity practices, Flexion is a business associate and requires a signed BAA with every practice/EHR partner before any pilot involving real patient data.

- Encryption in transit (TLS) and at rest for the session store; row-level access control scoped to a patient's assigned provider(s).

- Audit logging on every read/export of a patient's session data, and on every push to an EHR.

- Model/version pinning: because scoring thresholds directly affect what gets flagged to a clinician, every score must be traceable to the exact scoring model version that produced it (see ScoringModelVersion in Section 4).

- Device/browser camera permissions are the only patient-side access required — no separate app install is strictly necessary for the web version, which matters for a telehealth population with variable technical comfort.

# 7. Pilot Validation Protocol

The stated goal is concrete: does the AI's movement score agree with an in-person physical therapist's assessment of the same patient doing the same movement? This is a concurrent-validity study, not a usability study, and should be run before any clinical or billing claim is made. The literature review sharpens the priority order: movements already scored High confidence (squat depth, sit-to-stand timing, gait cadence) are the strongest early candidates for confirming product-specific validity quickly; Moderate and Exploratory movements (arm raise, balance) should be weighted more heavily in the pilot's sample size and analysis, since published evidence gives less basis for expecting agreement.

## 7.1 Design

- 3–5 telehealth providers treating chronic pain / MSK patients; target 30–50 patients for adequate statistical power on correlation and agreement estimates.

- For each patient's first cohort visit, run the Flexion routine and a licensed PT's in-person (or synchronous video-observed) assessment of the same movements, same day, in randomized order to avoid learning/fatigue bias.

- Repeat at 2–3 follow-up points over 6–8 weeks to also assess test–retest reliability and whether the AI tracks the same improvement/decline trend the PT documents.

- Include older adults and patients with mobility limitations explicitly in recruitment — the reviewed literature repeatedly flags insufficient testing in these populations as an open gap, and they are Flexion's target population for fall-risk and chronic-pain use cases.

## 7.2 Statistics

- Intraclass correlation coefficient (ICC) between AI and PT scores per movement — the same metric used in the published markerless-vs-marker-based validation literature, so results are comparable to prior work.

- Bland–Altman plots per movement to visualize bias and the range of disagreement, not just a single correlation number — mirroring the analysis used in the reviewed clinical-hallway markerless validation study.

- Sensitivity/specificity of the AI's flags (e.g. "limited_depth", "slow_sit_to_stand_time", "fall_risk_review") against the PT's clinical judgment of whether that finding was clinically significant.

## 7.3 Exit criteria

A movement's evidence tier in the product should be revised based on pilot results, not just literature precedent — a High-confidence tier earned from general markerless-capture literature can be downgraded if Flexion's own pilot disagrees with PT assessment, and an Exploratory tier (like balance) can be upgraded once product-specific validity is demonstrated. A movement should not ship to general clinical use, or claim a higher evidence tier, until its ICC and Bland–Altman results clear a threshold agreed with the pilot's clinical advisors — accuracy earns adoption; nothing else on the roadmap matters if this step is skipped.

# 8. Business Model Notes

| Segment | Entry point | Expansion |
| --- | --- | --- |
| Chronic pain / MSK telehealth practices | Between-visit functional monitoring, billed under RTM (98975–98981). | Progress tracking across a full care episode. |
| Senior health plans | Fall-risk scoring from the balance and sit-to-stand modules. | Population-level fall-risk dashboards. |
| Outpatient rehab / PT | Home exercise adherence + progress tracking between in-clinic visits. | Discharge-readiness scoring. |

Pricing in the $200–500/provider/month range (scaling with patient volume and EHR integration depth) is reasonable relative to RTM billing potential per patient per month, but the pilot's validity results should inform final pricing — a product that reliably reduces PT documentation time and supports RTM billing justifies the higher end; a screening-only tool without validated accuracy does not.

# 9. Immediate Next Steps

- Run the prototype with 2–3 internal testers across different cameras/lighting to sanity-check the scoring thresholds and framing-check behavior before any patient touches it.

- Recruit the 3–5 pilot providers and their clinical partners (PTs) for the validation study in Section 7, explicitly including older-adult and mobility-limited patients.

- Build the Session API + real auth, replacing the prototype's localStorage substitute, and stand up the FHIR Observation export end-to-end against one EHR sandbox (e.g. a SMART on FHIR sandbox) before the first real patient session.

- Draft and execute BAAs with pilot practices before any real patient data is captured.

# References

Core citations from the structured literature review (literature_research_on_the_app.docx, literature_citations_on_app.csv):

1. Chakka K, Wu R, Belko S, Annaswamy T. “Determining Gaps in Current Physiatric Tele-Physical Assessments via a Needs Assessment Survey.” American Journal of Physical Medicine & Rehabilitation, 2023. https://consensus.app/papers/determining-gaps-in-current-physiatric-telephysical-chakka-wu/29cafd2f88e85da7a9470a11f2a103a1/

2. Wade L, Needham L, McGuigan P, Bilzon J. “Applications and limitations of current markerless motion capture methods for clinical gait biomechanics.” PeerJ, 2022. https://consensus.app/papers/applications-and-limitations-of-current-markerless-wade-needham/99638d434aeb52bfb793fbcafd601174/

3. Zischke C, Simas V, Hing W, Milne N, Spittle A, Pope R. “The utility of physiotherapy assessments delivered by telehealth: A systematic review.” Journal of Global Health, 2021. https://consensus.app/papers/the-utility-of-physiotherapy-assessments-delivered-by-zischke-simas/9d164ce4d37b5b03ae6c92e8f6eba165/

4. Scott B, Blana D, Chadwick E, Philp F, Seyres M. “Healthcare applications of single camera markerless motion capture: a scoping review.” PeerJ, 2022. https://consensus.app/papers/healthcare-applications-of-single-camera-markerless-scott-blana/c74772b19a9a5852995bf30e78f0ffd9/

5. Scataglini S, Abts E, Van Bocxlaer C, Van den Bussche M, Meletani S, Truijen S. “Accuracy, Validity, and Reliability of Markerless Camera-Based 3D Motion Capture Systems versus Marker-Based 3D Motion Capture Systems in Gait Analysis: A Systematic Review and Meta-Analysis.” Sensors, 2024. https://consensus.app/papers/accuracy-validity-and-reliability-of-markerless-scataglini-abts/b2ea274d06605ac19019c898cf994898/

6. Horsak B, Eichmann A, Lauer K, Prock K, Krondorfer P, Siragy T, Dumphart B. “Concurrent validity of smartphone-based markerless motion capturing to quantify lower-limb joint kinematics in healthy and pathological gait.” Journal of Biomechanics, 2023. https://consensus.app/papers/concurrent-validity-of-smartphonebased-markerless-horsak-eichmann/e9c6da483a905903af8e7053f53a999e/

7. Pratapneni A, Halvorson R, Silvestros P, Harris N, Bailey J. “Validating Single-Camera Pose Estimation Against Multi-Camera Motion Capture for Accessible Biomechanical Assessment.” IEEE Access, 2026. https://consensus.app/papers/validating-singlecamera-pose-estimation-against-pratapneni-halvorson/37ceb9d8a3ea5aeaad215db374e1f570/

8. Ruder M, Malek M, Gafoor F, Di Bacco V, Outerleys J, Laende E. “Evaluating the Reliability of Markerless Motion Capture for Joint Angle Estimation in a Clinical Hallway Setting Compared to a Traditional Laboratory Setting.” Orthopaedic Proceedings, 2026. https://consensus.app/papers/evaluating-the-reliability-of-markerless-motion-capture-ruder-malek/5b1d256058dd5c90a78d644289464e1c/

9. Lee U, Lee S, Kim S-A, Kim Y, Lee S. “Validity and reliability of single camera markerless motion capture systems with RGB-D sensors for measuring shoulder range-of-motion: a systematic review.” Frontiers in Bioengineering and Biotechnology, 2025. https://consensus.app/papers/validity-and-reliability-of-single-camera-markerless-lee-lee/e51cfeae42e55f46870aa809f98c26e8/

10. Lam WWT, Tang Y, Fong K. “A systematic review of the applications of markerless motion capture (MMC) technology for clinical measurement in rehabilitation.” Journal of NeuroEngineering and Rehabilitation, 2023. https://consensus.app/papers/a-systematic-review-of-the-applications-of-markerless-lam-tang/6aa2b704d62c5fbb806b5963e50b50d1/

11. Johnson MA, Khilnani TK, Hyun A, Amen TB, Varady NH, Nwachukwu BU, Dines JS. “The State of Telemedicine, Telerehabilitation, and Virtual Care in Musculoskeletal Health: A Narrative Review.” HSS Journal, 2025. https://consensus.app/papers/the-state-of-telemedicine-telerehabilitation-and-virtual-johnson-khilnani/7c246c8b5ae25b18b1a00e20424f2a68/

12. Laskowski E, Johnson SE, Shelerud R, et al. “The Telemedicine Musculoskeletal Examination.” Mayo Clinic Proceedings, 2020. https://consensus.app/papers/the-telemedicine-musculoskeletal-examination-laskowski-johnson/fb0c1438fca15d0cba33da55c324f51f/

13. Tanaka MJ, Oh L, Martin S, Berkson E. “Telemedicine in the Era of COVID-19.” Journal of Bone and Joint Surgery, 2020. https://consensus.app/papers/telemedicine-in-the-era-of-covid19-tanaka-oh/6e967fa68f2f5f4aa15a7ab4ce2bf784/

14. Tenovi. “RTM CPT Codes 2026.” https://www.tenovi.com/rtm-cpt-codes-2026/

15. Coherent Market Insights. “Digital Musculoskeletal (MSK) Care Market Size 2026–2033.” https://www.coherentmarketinsights.com/industry-reports/digital-musculoskeletal-msk-care-market

Full list of 20 reviewed papers, abstracts, and evidence-coverage scoring is in literature_citations_on_app.csv and literature_research_on_the_app.docx.